## CT128
Bidders shall complete this Technical Compliance Matrix providing responses to all items in the Specifications.
Requirements |  |  | Bidder's Compliance (Y/N) | Bidder's Technical Response | Evidence of Technical Compliance Bidder shall state the proposal section(s), page number(s), and add comments as applicable
Technical requirements:
1 | Functional and performance requirements and scanning parameters | CT scanning system with at least 128 slices suitable to perform a wide range of scans in patients of all ages and a diverse range of diseases, with a special focus on lung diseases.
The CT scanner shall be suitable for all relevant cross-sectional imaging requirements in diagnostic radiology.
Number of actual slices per 360° rotation shall be not less than 128 slices at all speeds.
Reconstructed slice width options range from at least 0.5 mm to 10 mm.
rotation time ≤0.35 sec per 360° (faster times preferred)
Retrospective reconstruction shall be possible of raw data files with the capacity to modify parameters such as field of view (FOV).
At least the following scanning modes shall be possible: scan projection radiograph (SPR), axial and spiral.
The SPR length shall be at least 1500 mm and the minimum width 500 mm or less. It shall be possible to obtain the SPR from anteroposterior (AP) or posteroanterior (PA) or left to right or right to left directions.
Multiple volumetric studies capabilities.
The system shall have automated dose control and milliampere (mA) control software that automatically adjusts mA for patient size, adjusts mA along the z-axis, and modulates mA in the X-Y plane (angular modulation). Additionally, the system shall support 3D Dose Modulation (X-Y-Z axis modulation)
The CT scanner shall have conventional built-in lasers or light beams, which indicate the coincidence of the centre of rotation and scan position.
The system shall be interconnected (all the workstations, any laser systems, printers etc.) and the CT scanner shall be able to be networked at the site to allow transfer of CT data sets in Digital Imaging and Communications in Medicine (DICOM) format to a hospital information system (HIS), radiology information system (RIS) and picture archiving and communication system (PACS) system.
Dose computation and display: the system shall display CT dose index (CTDI) – volumetric (CTDIvol) and weighted (CTDIw) – and dose-length product (DLP) and have the capacity to transfer this information to the exam record.
Organ-Based Dose Reduction technology
Dose optimization tools shall be available.
Advanced dose reduction technology, including motion correction and iterative reconstruction, AI-based or equivalent
2 | Patient table | The CT scanner shall have a tabletop made of resistant and radiotransparent material.
Table with minimum dimensions of 200-250 cm x 40-50 cm
Vertical moving minimum range 44–100 cm
The speed of horizontal movement shall be variable with a maximum speed of at least 100 mm per sec.
The scannable range shall be at least 160 cm.
Maximum loading capacity of not less than 200 kg without any change in stated performance technical specifications (such as the positioning accuracy).
Capability of automatic patient isocentring (preferable).
3 | Gantry | Continuous rotating slip ring with umber of rows not less than 128
A gantry aperture of at least 70 cm.
Scan FOV of at least 50 cm.
Balancing system and laser lights to support centring.
Preferable: gantry tilt at least ± 30 degrees (preferable), controllable by the console; also (remote).
4 | X-ray system (tube and generator) and detectors | High-frequency generator with gantry implanted microprocessor.
The X-ray tube shall have dual focal spots. The contractor shall state the size of the focal spots.
Tube focal spots not higher than 0.6 - 1.2 mm
Power rating of at least 70 kW.
kVp range of at least 80 kV to 140 kV.
mA range of at least 100-800 mA
Maximum anode cooling rate of at least 1000 kHU/min.
Anode heat storage capacity of not less than 8 MHU.
High performance, low noise, high data density, active response data acquisition system.
Solid state detectors of width not more than 1 mm
At least 40 mm detector coverage
Automatic calibration
There shall be at least 700 elements per row and 128 or more detectors not interpolated
Quantum efficiency ≥ 99%
High discharging speed ≤ 0.5 ms preferred for high-speed imaging
5 | Hardware and software for operating console | 1. Raw data reconstruction hardware:
Latest generation system with high performance processors.
64 GB (Minimum), 128 GB+ (Recommended)
Raw data hard disk of at least 4TB
parallel multi-tasking level.
2. Double monitor acquisition (operation) console with at least the following characteristics:
Ergonomic console with height-adjustable stand & anti-fatigue design
Capability for simultaneous scanning and reconstruction.
multi-core high-performance CPU for real-time processing
At least the following functions shall be available: simultaneous scanning and routine analysis, simultaneous scanning and archiving and/or hard copying and simultaneous scanning and transfer to second console/workstation.
Last generation processor.
Two 24”-32” 4K UHD medical-grade monitors
Keyboard and mouse.
At least 32GB+ of RAM.
At least 2TB image storage hard disk.
CD/DVD driver with integrated CD reviewing software for DICOM stored images, USB and ability to cloud storage
The CT system shall be fully DICOM compliant. The DICOM shall support at least the following: DICOM 3.0 print, storage, send/receive, and query/retrieve.
DICOM data high transmission speed to and from the workstation.
Verbal/audio bi-directional communication system shall be provided between the operator and the patient.
At least local area network (LAN) connection driver.
UPS for the operator Console
Acquisition (operation) console’s software at least with the following functions:
Easy access software interface.
Positioning digital radiography with a length of at least 160 cm.
Initial definition of the exam’s protocol with real-time availability to modify it.
List of defined exam’s protocol.
Support for multiple scanning modes, including: Planner (Scout/Localizer), Axial (Step-and-Shoot), Helical (Spiral), Continuous Dynamic (Cine), Bolus Tracking, Multi-Phase Contrast, ECG-Gated Cardiac Imaging, Ultra-Low Dose, High-Pitch (Turbo Flash), Perfusion CT, Dual-Energy (Spectral), Metal Artifact Reduction (MAR), and Tumor Tracking.
For contrast-enhanced exams, scan synchronization with the arrival of the administered contrast (bolus test and bolus tracking).
Multi-perspective and cardiovascular examination software with at least the following requirements:
best reconstruction timing choice.
electrocardiogram (ECG) tracing and image-gating procedure.
automatic multi-phase reconstruction of cardiac kinetic function.
Software for low dose protocols included.
6 | Hardware and Software for independent workstation | Independent double monitor workstation with at least the following characteristics:
Ergonomic with height-adjustable stand & anti-fatigue design
High-performance workstation for post-processing and advanced clinical applications.
Latest generation processor.
Two ≥24” 4K UHD medical-grade monitors
Keyboard and mouse.
At least 64 GB+ of RAM.
Not less than 4 TB image storage . Type of Disk to be povided ( HDD, SSD , other ..) or equivalent cloud-based storage
CD/DVD drive with DICOM viewer software, with optional USB export capability
Super DVD writer: Standalone or integrated system with the CT scanner workstation.
The CT system shall be fully DICOM compliant. The DICOM shall support at least the following: DICOM 3.0 print, storage, send/receive, and query/retrieve.
DICOM data high transmission speed to and from the workstation.
At least LAN connection driver.
UPS for the Independant Workstation
Software features for the independent workstation described in the previous point:
3D Imaging & Reconstruction
MIP, MinIP, VRT, MPR, Surface 3D, Curve MPR, Advanced MPR with Thin-Slice Reconstruction
Dynamic & Contrast-Enhanced Scanning
· Dynamic Scan, Bolus-Tracking CT, Multi-Phase Contrast Imaging, Enhanced Synchronization with Scanner Movement
· Cardiovascular Imaging & Analysis
· Coronary Calcium Scoring, Plaque Characterization, ECG-Gated Image Reconstruction for Clearer Cardiac Imaging
Vascular & Angiographic Analysis
· 2D & 3D Vessel Structure Analysis, Higher-Resolution Vessel Segmentation for Blood Flow Analysis
Neuro Imaging & Stroke Analysis
· Digital Subtraction Angiography (DSA) for Neuro-CT, Perfusion CT Mapping (CBF, MTT) for Stroke Analysis
Lung Imaging & Pulmonary Analysis
· Full Lung Analysis Package, Dedicated Emphysema & Fibrosis Analysis, Pulmonary Function Evaluation
Oncology & Tumor Imaging
· Virtual Colonoscopy, Lung Nodule Detection, Improved Lesion Segmentation & Multi-Timepoint Tumor Tracking
Bone Imaging & Artifact Reduction
· Automatic Bone Subtraction, Metal Artifact Reduction (MAR), Bone Density Quantification for Orthopedic Imaging
Workflow & Post-Processing Enhancements
· Basic ECG-Gated Reconstruction, Improved ECG-Gated Perfusion CT Mapping, Multi-Scan Comparison Tools
· The system shall support automated filming and image processing, allowing seamless integration with DICOM-based storage and printing workflows.
Dose tracking and modulation control, enabling real-time monitoring of CTDIvol, DLP, and automatic dose modulation
All the software shall be provided for reporting procedure with DICOM, pdf, rtf format export availability to enable transfer of images for tele-radiology to electronic health records, to other medical facilities or to others as warranted.
7 | Image quality | The reconstruction matrix shall be at least 1024 x 1024 pixels
Display matrix not less than 1024 x 1024 pixel.
In plane spatial resolution ≤ 0.3 mm
Low-contrast detectability (or resolution) ≤ 3 mm at 0.3% contrast at ≤ 10 mGy
High-contrast spatial resolution: at least 18 lp/cm at 0% modulation transfer function (MTF).
Metal artefacts reduction algorithm
Reconstruction speed: Minimum 190 slices/sec or equivalent processing performance
Reconstruction speed, in both spiral and axial scan, with every artefact correction, at least 20 images/sec VRT.
Table and bone removal (image subtraction) function available.
Virtual endoscopic view available
8 | Power supply | Power input to be approximately 380 V, 50/60 Hz, 90 kVA, triphasic electrical source.
The provider will supply direct electric connection to the three-phase power supply network and the connection will be effectuated directly to the network with thermomagnetic disconnecting switch.
Protectors against power surge (over-voltage and over-current) line conditions.
UPS : online UPS with maintenance-free batteries for the backup of the entire system for at least 30 minutes.
9 | Accessories | Phantoms for Clibration
Lead Aprons (0.5mm Pb equivalent, full-body protection)
Thyroid Shields (adjustable, 0.5mm Pb equivalent)
Lead Glasses (protective eyewear against scatter radiation)
Lead Gloves (for operator hand protection in interventional procedures)
Patient Lead Shielding (gonadal)
Patient Handling & Comfort Accessories: Bidders to indicate what is offered
10 | CT Injector | Angiography designed
Electromechanical syringe
Electrical and automatic volume stop mechanism
Must accept syringe barrel heater
Mobile dual-head auto injector, compatible with the CT scanner’s bolus tracking system, with automatic flow rate adjustment. Must include at least 100 startup kits (syringe and tube
Any variation to be indicated in the offer.
11 | Regulatory Compliance | FDA or CE Electrical Safety Standard IEC 60601-IEC 60601-2-44 ISO 13485 - ISO 9001
IAEA and EPA Safety Standards
The equipment should comply with one of the energy efficiency standards and environmental / sustainability good international industry practices for example International Energy Efficient labels such as EE, EU EE, USA ENERGY STAR
The system must comply with all relevant accreditation bodies and local regulation.
12 | Site Preparation and Readiness
The bidder shall:
Conduct a site assessment to verify room suitability.
Ensure the floor load capacity supports the CT scanner
Assess and upgrade radiation shielding as necessary as per regulations for walls , doors , windows .....
Ensure proper ventilation to maintain room temperature as per manufacturer recommendations and provide split AC if required
Confirm compliance with all regulatory standards for safety and radiation protection.
Verify room layout and accessibility for equipment transport and workflow. Perform all site works related to the equipment Assess and upgrade the warning light outside the Gantry Room
13 | Training
10 days on-site training
Short Operating Procedure (SOP) in english and/or French as per the hospital requirement
14 | Deliverable Data Items
The Bidder shall provide two (2) complete sets of operation and servicing manuals and technical drawings in the English and French languages.
15 | Warranty
Warranty shall also cover hardware and software, including further upgrades and updates;
Warranty shall include all necessary spare parts, shipment to the Site, costs of replacement (including labour) and disposal of faulty parts.
Specific warranty inclusions and exclusions shall be clearly listed.
16 | Maintenance
The Bidder shall include full maintenance services during the standard warranty period.
Full maintenance services during the warranty period shall include:
- preventive maintenance;
- on-call interventions;
- any software update/upgrade for the System that will be become available;
- all necessary replacement and spare parts.
As part of the on-site acceptance, the Bidder shall provide to the local engineer and to the hospital medical physicist a plan for preventative maintenance and the name and contacts of a service representative/office for on-call maintenance intervention;
The Bidder shall ensure that a suitable qualified person can be onsite within twenty Four (24) hours following an unexpected breakdown and to solve the problem within the next twenty (48) hours during the warranty period.